Arrhythmia Detection:

INTRODUCTION:

Standard ECG speed = 25 mm/second

| Box Type | Time (seconds) | Time (ms) | Voltage (mV) |
| --- | --- | --- | --- |
| 1 small box | 0.04 s | 40 ms | 0.1 mV |
| 1 big box | 0.20 s | 200 ms | 0.5 mV |

5 small boxes = 1 big box

| Normal Ranges |  |  |
| --- | --- | --- |
| Category | Time (s) | Time (ms) |
| P wave | 0.06–0.11 seconds | 60–110 ms |
| PR interval | 0.12–0.20 seconds | 120–200 ms |
| QRS complex | 0.06–0.10 seconds | 60–100 ms |
| QT interval | 0.36–0.44 seconds | 360–440 ms |
| T wave | 0.10 – 0.25 seconds | 100 – 250 ms |

The arrhythmia detection algorithm typically includes the following key components:

Algorithm Flow for Arrhythmia Detection Using 2-Channel ECG

#### 1. Signal Acquisition

Input: Continuous ECG from a wearable biosensor with 2 leads (Lead I & II or vector equivalents).

Sampling Rate: Typically, 250–500 Hz.

#### 2. Preprocessing

Noise Filtering:

  - Bandpass filter (e.g., 0.5–40 Hz): Removes baseline wander & high-frequency noise.

  - Notch filter at 50/60 Hz: Removes powerline interference.

  - Adaptive filtering: Handles motion artifacts.

Signal Quality Index (SQI): To reject corrupted segments.

#### 3. QRS Complex Detection

Algorithms: Pan-Tompkins, Hamilton-Tompkins, or wavelet-based detection.

Output: RR intervals, beat locations.

#### 4. Feature Extraction (Per Beat or Per Window)

Time-domain features:

  - RR intervals, RR variability

  - QRS duration

  - PR and QT intervals

Morphological features:

  - P wave presence/absence

  - ST elevation/depression

  - T wave inversion

Derived features:

  - Heart rate

  - Irregularity indices

#### 5. Rhythm Classification

Rule-based heuristics (simple logic for initial detection):

AF: Irregular RR intervals, absence of P waves.

PVC: Wide QRS (>120 ms), no preceding P wave, compensatory pause.

PAC: Early beat with narrow QRS, abnormal P wave.

VT: Series of wide QRS complexes (>100 bpm), absence of P waves.

Machine Learning / Deep Learning:

Trained on labelled ECG segments.

CNN/LSTM models process raw signal or extracted features.

#### 6. Decision and Alerting

Thresholds for triggering alerts (e.g., sustained VT > 30 sec).

Display flagged ECG segments with time stamps.

Option to push to cloud or physician dashboard.

The following are the steps involved for analysing an ECG rhythm,

- Step 1: Calculate heart rate.

- Step 2: Determine regularity.

- Step 3: Assess the P waves.

- Step 4: Determine PR interval

- Step 5: Determine QRS duration.

STEP 1: Calculate Heart Rate:

There are three methods for calculating the Heart rate in ECG,

- Six Seconds method

- Rule of 1500

- Sequence method

6 Seconds Method:

The six-second method is a simple and practical approach for estimating heart rate from an electrocardiogram (ECG) tracing, especially when dealing with irregular rhythms or when a rapid assessment is needed. Here's how it works:

- Locate a Starting Point: Choose a point on the ECG tracing, such as the beginning of a particular complex (e.g., the R wave of a QRS complex), to start your count.

- Count the Number of R Waves: Count the number of R waves that occur within a six-second interval on the ECG tracing. You can either use a timer to precisely measure six seconds or count the number of large squares on the ECG paper (each large square typically represents 0.2 seconds).

- Calculate Heart Rate: Multiply the number of R waves counted by 10 to estimate the heart rate in beats per minute (bpm). This is because six seconds is one-tenth of a minute (60 seconds).

For example, if you count 11 R waves within a six-second interval, the estimated heart rate would be 11x 10 = 110 bpm.

The six-second method provides a quick and approximate assessment of heart rate, which can be useful in clinical settings for triage, initial evaluation, or monitoring changes in heart rate over time.

1500 Method:

The 1500 method is a simple technique used to estimate heart rate from an electrocardiogram (ECG) tracing, particularly when measuring heart rate over a short duration or when precise measurements are not necessary. Here's how it works:

- Identify R Waves: Locate the R waves on the ECG tracing. The R wave represents ventricular depolarization, marking the beginning of each cardiac cycle.

- Count the Number of Small Squares: Count the number of small squares (each representing 0.04 seconds or 40 milliseconds) between consecutive R waves.

- Apply the Formula: Divide 1500 by the number of small squares counted to estimate the heart rate in beats per minute (bpm). This formula is derived from the fact that there are 1500 small squares in a minute (60 seconds).

For example, if there are 30 small boxes between consecutive R waves, the estimated heart rate would be 1500 / 30 = 50 bpm.

The 1500 method provides a quick approximation of heart rate, especially in situations where rapid assessment is required or when dealing with irregular rhythms.

Sequence Method:

The sequence method, also known as the "300-150-100-75-60" method, is another approach to quickly estimate heart rate from an ECG tracing, particularly when calculating heart rate over a short period or when faced with irregular rhythms. Here's how it works:

- Identify the R waves: Locate a prominent R wave on the ECG tracing. The R wave corresponds to ventricular depolarization, indicating the beginning of a cardiac cycle.

- Count the number of large squares between R waves: Each large square on standard ECG paper typically represents 0.2 seconds (or 200 milliseconds). Count the number of large squares between consecutive R waves.

- Apply the conversion formula: Use the formula "300 divided by the number of large squares" to estimate heart rate. This formula is based on the fact that there are 300 large squares in a minute (60 seconds).

  - If there is 1 large square between R waves, the heart rate is approximately 300 bpm.

  - If there are 2 large squares between R waves, the heart rate is approximately 150 bpm.

  - If there are 3 large squares between R waves, the heart rate is approximately 100 bpm.

  - If there are 4 large squares between R waves, the heart rate is approximately 75 bpm.

  - If there are 5 large squares between R waves, the heart rate is approximately 60 bpm.

This method provides a quick estimation of heart rate during situations where precise measurement may not be possible or practical, helping healthcare professionals to promptly assess patients' cardiac status and make informed decisions.

STEP 2: Determine Regularity:

As part of rhythm assessment, it is important to determine whether the ECG trace is regular or irregular.

- Carefully measure from the tip of one R wave to the next, from the beginning to the end of the tracing.

- A rhythm is considered “regular or constant” when the distance apart is either the same or varies by 1 ½ small boxes or less from one R wave to the next R wave.

STEP 3: Assess P waves:

- Is there a single P wave before each QRS complex?

- Do all-P waves look alike?

- Measure the P wave duration

- Single P Waves before Each QRS Complex: Start by examining the relationship between the P waves and QRS complexes. Ideally, there should be a single P wave preceding each QRS complex. This indicates normal atrioventricular conduction, where the electrical impulse originates in the sinoatrial (SA) node, travels through the atria (producing the P wave), and then activates the ventricles (producing the QRS complex).

- Uniformity of P Waves: Next, evaluate whether all the P waves look alike throughout the ECG tracing. Normal P waves should have consistent morphology (shape) and amplitude. If there are variations in the appearance of the P waves, it could suggest abnormalities such as ectopic atrial rhythms or atrial enlargement.

- Measurement of P Wave Duration: Measure the duration of the P waves. The duration is typically assessed from the beginning to the end of the P wave where it returns to the baseline. Normal P waves have duration of less than 0.12 seconds (120 milliseconds). If the P wave duration exceeds this threshold, it may indicate atrial enlargement or conduction abnormalities.

Common P Wave Abnormalities

- P mitrale (bifid P waves), seen with left atrial enlargement.

- P pulmonale (peaked P waves), seen with right atrial enlargement.

- P wave inversion, seen with ectopic atrial and junctional rhythms.

- Variable P wave morphology, seen in multifocal atrial rhythms.

STEP 4: Assess PR Interval:

The PR interval represents the time it takes for the electrical impulse to travel from the atria to the ventricles, passing through the atrioventricular (AV) node.

Key Considerations:

- Normal vs. Abnormal (Short/Prolonged) PR Interval: A normal PR interval typically ranges from 0.12 to 0.20 seconds (120 to 200 milliseconds) on the ECG. Deviations from this range may indicate abnormal AV conduction. A prolonged PR interval suggests delayed AV conduction, while a short PR interval indicates accelerated conduction.

Normal PR Interval:

- A normal PR interval typically falls within the range of 0.12 to 0.20 seconds (120 to 200 milliseconds) on the ECG.

- It indicates normal atrioventricular (AV) conduction, where the electrical impulse travels from the atria to the ventricles in a timely manner through the AV node.

- A normal PR interval suggests coordinated depolarization of the atria and ventricles.

Prolonged PR Interval:

- A prolonged PR interval is defined as exceeding 0.20 seconds (200 milliseconds) on the ECG.

- It indicates delayed atrioventricular conduction, which can occur due to various factors such as first-degree heart block, second-degree heart block (Type I or Type II), or medications affecting AV conduction.

- Prolonged PR intervals may lead to delayed ventricular activation and subsequent heart rhythm disturbances.

Short PR Interval:

- A short PR interval is characterized by duration less than 0.12 seconds (120 milliseconds) on the ECG.

- It suggests accelerated atrioventricular conduction, where the electrical impulse travels rapidly from the atria to the ventricles.

- Short PR intervals may be seen in conditions such as Wolff-Parkinson-White (WPW) syndrome, junctional rhythms, or pre-excitation syndromes.

- Consistency of PR Interval:

- After determining whether the PR interval is normal, prolonged, or short, assess its consistency throughout the ECG tracing. A constant PR interval indicates stable atrioventricular conduction and rhythm.

- In contrast, a changing PR interval, where the duration varies from beat to beat, may suggest underlying heart block, atrial fibrillation, or other arrhythmias affecting AV conduction.

STEP 5: Assess QRS Complex:

The QRS complex is a distinct waveform on the ECG that represents ventricular depolarization, or the electrical activity associated with the contraction of the heart's ventricles. It typically consists of three deflections: Q, R, and S waves.

Arrhythmia: An arrhythmia is an abnormality in the rate, rhythm, or sequence of the heartbeat due to a disturbance in the heart's electrical conduction system

Sinus Rhythm:

Sinus rhythm refers to the normal heart rhythm originating from the sinoatrial (SA) node, the heart's natural pacemaker. It is characterized by regular depolarization of the atria followed by ventricular depolarization, resulting in coordinated atrial and ventricular contractions.

| Feature | Characteristics |
| --- | --- |
| P wave Amplitude (mV / mm) | 0.1–0.3 mV (1–3 mm) |
| P wave Width (ms) | ≤ 120 ms (usually 80–110 ms) |
| PR Interval (ms) | 120–200 ms |
| QRS Amplitude (mV / mm) | Varies by lead, but typically 0.5–2.5 mV (5–25 mm) |
| QRS Width (ms) | ≤ 100 ms (normal is 80–100 ms) |
| ST Segment Deviation (mV / mm) | Isoelectric (no significant deviation: <1 mm or <0.1 mV) |
| T wave Amplitude (mV / mm) | 0.1–0.5 mV (1–5 mm), upright in most leads |
| QT Interval | 350–440 ms (corrected QTc < 440 ms for men, < 460 ms for women) |
| Regularity | Regular (consistent RR intervals) |
| Rate (bpm) | 60–100 bpm |
| P wave Morphology | Upright in leads I, II, aVF; inverted in aVR; uniform shape |
| Other Notable Features | Each P wave followed by a QRS; 1:1 AV conduction; normal axis |

Sinus Tachycardia:

Sinus tachycardia is a cardiac rhythm characterized by a fast heart rate originating from the sinoatrial (SA) node. It is defined by a heart rate exceeding 100 beats per minute (bpm). Sinus tachycardia is typically a physiological response to factors such as exercise, stress, fever, pain, or anxiety. It can also occur due to medical conditions such as dehydration, hypovolemia, hyperthyroidism, or heart failure.

| Feature | Characteristics |
| --- | --- |
| P wave Amplitude (mV / mm) | ≤ 0.25 / 2.5 |
| P wave Width (ms) | < 120 |
| PR Interval (ms) | 120-200 |
| QRS Amplitude (mV / mm) | 0.5–1.0 / 5–10 |
| QRS Width (ms) | < 120 |
| ST Segment Deviation (mV / mm) | Isoelectric |
| T wave Amplitude (mV / mm) | ≤ 0.5 / 5 |
| QT Interval | Normal |
| Regularity | Regular |
| Rate (bpm) | > 100 |
| P wave Morphology | Normal, upright (may merge with preceding T) |
| Other Notable Features | Fast heart rate. |

Sinus Bradycardia

Sinus bradycardia is a cardiac rhythm characterized by a slow heart rate originating from the sinoatrial (SA) node. It is defined by a heart rate less than 60 beats per minute (HR < 60 BPM). Sinus bradycardia can be a normal finding in healthy individuals, particularly athletes or during sleep. In some cases, sinus bradycardia may indicate underlying heart conditions such as sick sinus syndrome, atrioventricular (AV) block, or myocardial infarction.

| Feature | Characteristics |
| --- | --- |
| P wave Amplitude (mV / mm) | ≤ 0.25 / 2.5 |
| P wave Width (ms) | < 120 |
| PR Interval (ms) | 120-200 |
| QRS Amplitude (mV / mm) | 0.5–1.0 / 5–10 |
| QRS Width (ms) | < 100 |
| ST Segment Deviation (mV / mm) | Isoelectric |
| T wave Amplitude (mV / mm) | ≤ 0.5 / 5 |
| QT Interval | Normal |
| Regularity | Regular |
| Rate (bpm) | < 60 |
| P wave Morphology | Normal, upright |
| Other Notable Features | Slow heart rate. |

Sinus Arrhythmia:

Sinus arrhythmia is a commonly encountered variation of normal sinus rhythm. Sinus arrhythmia characteristically presents with an irregular rate in which the variation in the R-R interval is more than 0.12 seconds (120 milliseconds). One type of sinus arrhythmia, called respiratory sinus arrhythmia, is when the heartbeat changes pace when you inhale and exhale. In other words, your heartbeat cycles with your breath. When you breathe in, your heart rate increases. When you exhale, it falls.

| Feature | Characteristics |
| --- | --- |
| P wave Amplitude (mV / mm) | ≤ 0.25 / 2.5 |
| P wave Width (ms) | < 120 |
| PR Interval (ms) | 120-200 |
| QRS Amplitude (mV / mm) | 0.5–1.0 / 5–10 |
| QRS Width (ms) | < 100 |
| ST Segment Deviation (mV / mm) | Isoelectric |
| T wave Amplitude (mV / mm) | ≤ 0.5 / 5 |
| QT Interval | Normal |
| Regularity | Irregular (cyclical with respiration) |
| Rate (bpm) | 60-100 |
| P wave Morphology | Normal, upright (morphology may slightly vary) |
| Other Notable Features | Irregular rhythm is related to breathing. |

Atrial Ectopic/Rhythms:

Underlying Rhythm

In ECG interpretation, the "underlying rhythm" refers to the fundamental, consistent rhythm of the heart, even when there are occasional, transient abnormalities or extra beats disrupting the overall regularity. It's the rhythm that the heart would be in if those isolated disturbances were absent.

Understanding the underlying rhythm guides treatment. For instance, treatment for a life-threatening ventricular tachycardia is very different from managing occasional benign premature atrial contractions on an underlying sinus rhythm.

PREMATURE ATRIAL CONTRACTION (PAC)

PACs are ectopic beats that occur before the expected time. A compensatory pause should follow the premature beat. The morphology of the P wave changes from a normal sinus beat. The QRS complex remains normal. When the premature or early signal triggers the heart to contract, there may not be much blood in the heart at that moment. This means there's not much blood to pump out. A pause and a strong beat may follow the extra heartbeat, making it feel like a skipped beat.

| Feature | Characteristics |
| --- | --- |
| P wave Amplitude (mV / mm) | Variable |
| P wave Width (ms) | Variable |
| PR Interval (ms) | Variable |
| QRS Amplitude (mV / mm) | 0.5–1.0 / 5–10 |
| QRS Width (ms) | < 100 |
| ST Segment Deviation (mV / mm) | Isoelectric |
| T wave Amplitude (mV / mm) | Variable |
| QT Interval | Normal |
| Regularity | Premature |
| Rate (bpm) | Underlying rhythm |
| P wave Morphology | Different morphology, early |
| Other Notable Features | Often followed by a non-compensatory pause. |

TYPES OF PAC TERMINOLOGIES:

- Atrial Pair/couplet: 2 PACs comes together

| Feature | Characteristics |
| --- | --- |
| P wave Amplitude (mV / mm) | Variable |
| P wave Width (ms) | Variable |
| PR Interval (ms) | Variable |
| QRS Amplitude (mV / mm) | 0.5–1.0 / 5–10 |
| QRS Width (ms) | < 100 |
| ST Segment Deviation (mV / mm) | Isoelectric |
| T wave Amplitude (mV / mm) | Variable |
| QT Interval | Normal |
| Regularity | Premature |
| Rate (bpm) | Underlying rhythm |
| P wave Morphology | Two premature P' waves |
| Other Notable Features | Two PACs in a row. |

- Atrial Run: 3 to 5 PACs comes together

| Feature | Characteristics |
| --- | --- |
| P wave Amplitude (mV / mm) | Variable |
| P wave Width (ms) | Variable |
| PR Interval (ms) | Variable |
| QRS Amplitude (mV / mm) | 0.5–1.0 / 5–10 |
| QRS Width (ms) | < 100 |
| ST Segment Deviation (mV / mm) | Isoelectric |
| T wave Amplitude (mV / mm) | Variable |
| QT Interval | Normal |
| Regularity | Rapid, regular or irregular |
| Rate (bpm) | > 100 |
| P wave Morphology | Rapid succession of premature P' waves |
| Other Notable Features | Three or more PACs in a row. |

- Atrial bigeminy: Every 2nd beat is PAC

| Feature | Expected Value / Change |
| --- | --- |
| P wave Amplitude (mV / mm) | Sinus: 0.1–0.3 mV; PAC: 0.05–0.25 mV (often smaller or biphasic) |
| P wave Width (ms) | Sinus: 80–110 ms; PAC: 60–100 ms (may be slightly shorter or longer) |
| PR Interval (ms) | Sinus: 120–200 ms; PAC: 80–160 ms (shorter due to early atrial firing) |
| QRS Amplitude (mV / mm) | Normal: 0.5–2.5 mV; remains unchanged unless aberrant conduction |
| QRS Width (ms) | Normal: 80–100 ms; may be slightly widened (~100–120 ms) if PAC occurs very early |
| ST Segment Deviation (mV / mm) | Unchanged or <0.1 mV unless underlying pathology |
| T wave Amplitude (mV / mm) | 0.1–0.5 mV; may appear altered after PACs |
| QT Interval | Appears irregular; QTc usually < 440 ms men, < 460 ms women |
| Regularity | Alternates: e.g., 800 ms (normal), 600 ms (PAC), repeat — irregular |
| Rate (bpm) | 50–90 bpm (depends on PAC frequency and whether all are conducted) |

- Atrial Trigeminy: Every 3rd beat is PAC

| Feature | Expected Value / Change |
| --- | --- |
| P wave Amplitude | Sinus: 0.1–0.3 mV; PAC: 0.05–0.25 mV (altered shape/axis) |
| P wave Width | Sinus: 80–110 ms; PAC: 60–100 ms |
| PR Interval | Sinus: 120–200 ms; PAC: 80–160 ms |
| QRS Amplitude | 0.5–2.5 mV; usually unchanged |
| QRS Width | 80–100 ms; possibly widened if PAC is very early |
| ST Segment Deviation | < 0.1 mV |
| T wave Amplitude | 0.1–0.5 mV |
| QT Interval | Variable; corrected QTc usually normal |
| RR Interval | Pattern: e.g., 800 ms (normal), 800 ms (normal), 600 ms (PAC), then repeat — regularly irregular |
| Rate | 60–90 bpm (depends on underlying rhythm and PAC frequency) |

ATRIAL ESCAPE BEAT (AEB):

An atrial escape beat is a delayed atrial impulse arising from an ectopic atrial focus, occurring when the sinus node fails to fire on time (sinus pause, bradycardia, or arrest). It acts as a backup to maintain the heart rhythm.

- Occurs late, not early like PACs.

- Different P wave morphology from sinus beats.

- PR interval is usually within normal limits.

- QRS is typically narrow, since impulse still travels through normal conduction pathways.

- No compensatory pause after the beat (unlike PVCs).

- Seen after a sinus pause, sinus arrest, or severe bradycardia.

| Feature | Characteristics |
| --- | --- |
| P wave Amplitude | 0.05–0.3 mV; may be lower amplitude or altered shape |
| P wave Width | 70–120 ms (may be slightly different than sinus P wave) |
| PR Interval | 120–200 ms (may be slightly longer or shorter) |
| QRS Amplitude | Normal, unless conduction abnormality present |
| QRS Width | Normal (80–100 ms) |
| ST Segment | Typically, unchanged unless underlying pathology exists |
| T wave Amplitude | Usually normal; depends on preceding QRS and repolarization |
| QT Interval | Unchanged unless bradycardia-induced QT prolongation |
| RR Interval | One prolonged RR interval (pause) before escape beat |
| Rate | Escape beat occurs after a pause (e.g., >1600–2000 ms RR) |
| P wave Morphology | Different axis or shape compared to sinus P wave |
| Other Notable Features | Single beat following sinus pause; no compensatory pause |

Blocked PAC

A blocked PAC (Premature Atrial Contraction) is an early atrial beat that fails to conduct to the ventricles, meaning no QRS complex follows the premature P wave.

A non-conducted PAC is not a sign of AV block pathology. It's a timing issue, not a conduction disease — the AV node just didn't have enough time to recover after the previous beat.

| Feature | Characteristics |
| --- | --- |
| P wave Amplitude | Variable |
| P wave Width | Variable |
| PR Interval | May be absent or very short (due to block) |
| QRS Amplitude | Absent for blocked PAC (since no ventricular conduction) |
| QRS Width | N/A (no QRS after blocked PAC) |
| ST Segment Deviation | N/A |
| T wave Amplitude | N/A for blocked PAC |
| QT Interval | Normal for conducted beats |
| Regularity | Irregular due to premature and blocked beats |
| Rate (bpm) | Reflects underlying rhythm |
| P wave Morphology | Abnormal (P') – two premature P waves (PACs) with different morphology |
| Other Notable Features | One or both PACs could be blocked, resulting in a missing QRS |

Paroxysmal Supraventricular Tachycardia (PSVT)

Paroxysmal Supraventricular Tachycardia (PSVT) is a sudden-onset, rapid heart rhythm that originates above the ventricles (i.e., in the atria or AV node) and typically causes a regular, narrow-complex tachycardia.

- Rapid, regular rhythm (~180 bpm)

- Narrow QRS

- No clear P waves or inverted retrograde P waves just after QRS

| Feature | Characteristics |
| --- | --- |
| P wave Amplitude (mV / mm) | Often not visible; if present, variable and may be retrograde |
| P wave Width (ms) | Variable or absent |
| PR Interval (ms) | Shortened or not measurable (due to hidden P waves) |
| QRS Amplitude (mV / mm) | 0.5–1.0 / 5–10 (normal range) |
| QRS Width (ms) | < 120 (narrow complex) |
| ST Segment Deviation (mV / mm) | Typically, isoelectric |
| T wave Amplitude (mV / mm) | Normal or slightly altered due to rate |
| QT Interval | Usually normal |
| Regularity | Regular |
| Rate (bpm) | 150–250 bpm |
| P wave Morphology | Absent or retrograde; different from sinus P waves |
| Other Notable Features | Sudden onset and termination |

Atrial Fibrillation (AF)

ATRIAL FIBRILLATION:

Atrial fibrillation (AF) refers to irregular atrial contraction, caused by chaotic impulses.

AF is an extremely common tachyarrhythmia that occurs due to multiple waves of spontaneous depolarization leading to fragmentation of the normal electrical activity. This causes muscle fibers to contract independently, leading to ‘fibrillation’ of the atrial muscle.

As multiple waves of electrical activity (varying strength) reach the AVN, it intermittently conducts these to the ventricles. This results in an irregular rhythm, which is characterized by three main ECG features:

- Irregularly irregular rhythm

- Absence of P waves (no coordinated atrial activity)

- Irregular, fibrillating baseline.

- Atrial rate: 300–600 bpm (fibrillatory waves)

- May be paroxysmal, persistent, long-standing persistent, or permanent

| Feature | Typical Values / Characteristics |
| --- | --- |
| P wave Amplitude (mV / mm) | Absent — replaced by fibrillatory (f) waves of <0.1–0.2 mV (fine) or >0.2 mV (coarse) |
| P wave Width (ms) | Not measurable — no organized P waves present |
| PR Interval (ms) | Not applicable |
| QRS Amplitude (mV / mm) | Normal: 0.5–2.5 mV |
| QRS Width (ms) | Normal: <120 ms, unless underlying conduction defect present |
| ST Segment Deviation (mV / mm) | Variable; <1 mm deviation possible due to irregular conduction |
| T wave Amplitude (mV / mm) | Variable; often normal but can be distorted due to irregular ventricular activation |
| QT Interval (ms) | Variable; QTc should be interpreted with caution due to variable RR intervals; typically, <450 ms |
| RR Interval (ms) | Highly irregular; variation >60 ms between beats common |
| Rate (bpm) | Ventricular rate: 90–170 bpm (can be lower or higher depending on AV nodal conduction) |
| P wave Morphology | Absent; replaced by erratic baseline fibrillatory activity |
| Regularity | Irregularly irregular RR intervals — hallmark of AF |
| Other Notable Features | No consistent P-QRS-T pattern; fibrillatory waves best seen in leads V1, II; may be fine or coarse |

Atrial Flutter:

In atrial flutter, the atrial heart rate typically ranges between 250 and 350 beats per minute, while the ventricular rate may vary. The rhythm can exhibit either irregularity or regularity, and it is characterized by distinctive saw-toothed P waves referred to as Flutter waves.

Atrial flutter is like atrial fibrillation, a common disorder that causes the heart to beat in irregular patterns. People with atrial flutter have a heart rhythm that's more organized and less chaotic than that of atrial fibrillation. Sometimes a person may have episodes of both atrial flutter and atrial fibrillation.

- Atrial rate: typically, 250–350 bpm

- Ventricular rate: often regular, depending on AV block (e.g., 2:1, 3:1)

- F waves: Sawtooth flutter waves, no isoelectric baseline

- QRS: usually narrow

- Can be regular or irregular

| Feature | Typical Values / Characteristics |
| --- | --- |
| P wave Amplitude (mV / mm) | Absent as distinct P waves; replaced by flutter ("F") waves of amplitude 0.2–0.4 mV |
| P wave Width (ms) | Not applicable — flutter waves typically appear as regular sawtooth waves |
| PR Interval (ms) | Not measurable — no true P waves; variable AV conduction (e.g., 2:1, 3:1) |
| QRS Amplitude (mV / mm) | Normal: 0.5–2.5 mV |
| QRS Width (ms) | Normal: <120 ms |
| ST Segment Deviation (mV / mm) | May show minor deviation, usually <1 mm, due to rapid rate |
| T wave Amplitude (mV / mm) | May be obscured or merged with flutter waves |
| QT Interval (ms) | Difficult to measure accurately due to flutter waves; QTc usually <450 ms |
| RR Interval (ms) | Regular if fixed AV block (e.g., 2:1), irregular if variable block |
| Rate (bpm) | Atrial rate: 240–340 bpm; Ventricular rate: depends on block (e.g., ~150 bpm for 2:1) |
| P wave Morphology | Not present; instead, "sawtooth" flutter waves best seen in leads II, III, aVF, V1 |
| Regularity | Regular ventricular response in fixed block (e.g., 2:1); irregular if block varies |
| Other Notable Features | Flutter waves (“F” waves) occur at ~300 bpm; AV conduction determines ventricular response; may mimic SVT |

SUPRAVENTRICULAR TACHYCARDIA:

Supraventricular tachycardia is a rapid heartbeat that develops when the normal electrical impulses of the heart are disrupted.

The heart rate in this case ranges from 180 to 250 beats per minute, indicating a very regular rhythm. The P wave is buried within the T wave, making it unidentifiable, and the QRS complex remains narrow. SVT is also known as regularly regular rhythm or narrow QRS complex tachycardia.

- Regular rhythm

- Narrow QRS (<120 ms) unless aberrantly conducted

- Sudden onset and termination

- P waves: often hidden in or just after QRS

| Feature | Typical Values / Characteristics |
| --- | --- |
| P wave Amplitude (mV / mm) | Often low: < 0.15 mV or may be absent |
| P wave Width (ms) | If visible: 80–110 ms |
| PR Interval (ms) | <120 ms or not measurable (P wave hidden in QRS or T wave) |
| QRS Amplitude (mV / mm) | Normal: 0.5–2.5 mV |
| QRS Width (ms) | <120 ms (narrow); may be slightly wider with rate-related aberrancy |
| ST Segment Deviation (mV / mm) | <1 mm mild ST depression may be seen due to rapid rate |
| T wave Amplitude (mV / mm) | Variable: ±0.1–0.3 mV; may merge with retrograde P waves |
| QT Interval (ms) | Appears short due to tachycardia; corrected QT (QTc) usually <450 ms |
| RR Interval (ms) | 240–400 ms (corresponds to 150–250 bpm) |
| Rate (bpm) | 150–250 bpm |
| P wave Morphology | Inverted in inferior leads (II, III, aVF) if retrograde; or absent |
| Regularity | RR interval regular: variation <10 ms between consecutive beats |
| Other Notable Features | Sudden onset/offset; no clear P-QRS relationship; response to vagal maneuvers or adenosine |

JUNCTIONAL ECTOPICS/ARRHYTHMIAS

Junctional rhythms are cardiac arrhythmias characterized by the initiation of electrical impulses from the atrioventricular (AV) junction, typically the atrioventricular node (AV node) or the junctional tissue near the AV node. In junctional rhythms, the SA node fails to initiate the heart's electrical activity, leading to the AV junction taking over as the heart's primary pacemaker.

Characteristics of junctional Beats:

  - P wave: Absent/ Inverted/ Retrograde

  - Rhythm: Regular

  - QRS: Normal

Premature Junctional Contraction (PJCs)

- PJCs are ectopic beats that occur before the expected time.

- A compensatory pause should follow the premature beat.

- The morphology of the P wave is absent, inverted, or retrograde P waves

- The QRS complex remains normal.

| Feature | Characteristics |
| --- | --- |
| P wave Amplitude (mV / mm) | Absent or inverted before QRS (short PR) or after |
| P wave Width (ms) | Variable |
| PR Interval (ms) | < 120 (if before) or N/A |
| QRS Amplitude (mV / mm) | 0.5–1.0 / 5–10 |
| QRS Width (ms) | < 100 |
| ST Segment Deviation (mV / mm) | Isoelectric |
| T wave Amplitude (mV / mm) | Variable |
| QT Interval | Normal |
| Regularity | Premature |
| Rate (bpm) | Underlying rhythm |
| P wave Morphology | Absent, inverted, or after QRS |
| Other Notable Features | Premature beat originating from the AV junction. |

JUNCTIONAL ESCAPE BEATS (JEBs)

- Beat occurs after the expected time.

- No compensatory pause after the escape beat.

- P-wave morphology - absent, inverted, or retrograde P waves

- QRS complex: Normal

| Feature | Characteristic |
| --- | --- |
| P wave Amplitude (mV / mm) | Absent or inverted before QRS (short PR) or after |
| P wave Width (ms) | Variable |
| PR Interval (ms) | < 120 (if present before QRS) or N/A |
| QRS Amplitude (mV / mm) | 0.5–1.0 / 5–10 |
| QRS Width (ms) | < 100 |
| ST Segment Deviation (mV / mm) | Isoelectric |
| T wave Amplitude (mV / mm) | Variable |
| QT Interval | Normal |
| Regularity | Occurs after a pause (not premature) |
| Rate (bpm) | Follows underlying rhythm; typically, 40–60 bpm if junctional rhythm takes over |
| P wave Morphology | Absent, inverted, or occurs after QRS due to retrograde atrial activation |
| Other Notable Features | Escape beat arising from the AV junction; occurs when SA node fails to fire or there's a long pause |

JUNCTIONAL RHYTHM:

This is an abnormal heart rhythm that originates from the atrioventricular (AV) junction, rather than the usual pacemaker of the heart-the sinoatrial (SA) node. It typically occurs when the SA node fails or is suppressed, prompting the AV node to take over.

- Heart rate: Usually between less than beats per minute (<60bpm).

- ECG features: Absent or inverted P waves, narrow QRS complexes.

Example: HR 50 bpm with inverted P wave

| Feature | Characteristic |
| --- | --- |
| P wave Amplitude (mV / mm) | Absent or inverted before QRS (short PR) or after |
| P wave Width (ms) | Variable |
| PR Interval (ms) | < 120 (if present before QRS) or N/A |
| QRS Amplitude (mV / mm) | 0.5–1.0 / 5–10 |
| QRS Width (ms) | < 100 |
| ST Segment Deviation (mV / mm) | Isoelectric |
| T wave Amplitude (mV / mm) | Variable |
| QT Interval | Normal |
| Regularity | Regular rhythm |
| Rate (bpm) | Typically 40–60 bpm (intrinsic rate of the AV junction) |
| P wave Morphology | Absent, inverted, or after QRS due to retrograde atrial activation |
| Other Notable Features | Rhythm originates from the AV junction; often seen when SA node is suppressed or non-functional |

ACCELERATED JUNCTIONAL RHYTHM:

Accelerated junctional rhythm is a rhythm arising from the AV junction with a heart rate typically between 60 and 100 beats per minute, which is faster than a normal junctional escape rhythm (which is usually 40–60 bpm), but slower than junctional tachycardia (which is usually >100 bpm).

- Regular rhythm with a rate of 60–100 bpm

- P waves may be absent, inverted (retrograde), or appear before, during, or after the QRS comple

Ex: HR 80 bpm with absent of P wave

| Feature | Characteristic |
| --- | --- |
| P wave Amplitude (mV / mm) | Absent or inverted before QRS (short PR) or after |
| P wave Width (ms) | Variable |
| PR Interval (ms) | < 120 (if present before QRS) or N/A |
| QRS Amplitude (mV / mm) | 0.5–1.0 / 5–10 |
| QRS Width (ms) | < 100 |
| ST Segment Deviation (mV / mm) | Isoelectric |
| T wave Amplitude (mV / mm) | Variable |
| QT Interval | Normal |
| Regularity | Regular rhythm |
| Rate (bpm) | 60–100 bpm (faster than junctional rhythm, but below tachycardia threshold) |
| P wave Morphology | Absent, inverted, or after QRS due to retrograde conduction |
| Other Notable Features | Often due to enhanced automaticity or digoxin toxicity; not typically seen in healthy individuals |

JUNCTIONAL TACHYCARDIA:

Junctional tachycardia is a supraventricular tachycardia characterized by a regular rhythm that originates from the AV junction with a rate typically between 100 and 180 bpm.

- Regular rhythm with rate >100 bpm.

- P waves: May be absent, inverted (retrograde), or occur before, during, or after the QRS.

- QRS complexes are usually narrow (<120 ms) unless there is aberrant conduction.

- Short PR interval if P waves precede the QRS.

| Feature | Characteristic |
| --- | --- |
| P wave Amplitude (mV / mm) | Absent or inverted before QRS (short PR) or after |
| P wave Width (ms) | Variable |
| PR Interval (ms) | < 120 (if present before QRS) or N/A |
| QRS Amplitude (mV / mm) | 0.5–1.0 / 5–10 |
| QRS Width (ms) | < 100 |
| ST Segment Deviation (mV / mm) | Isoelectric or may show subtle changes due to rate |
| T wave Amplitude (mV / mm) | Variable |
| QT Interval | Normal or slightly shortened due to high rate |
| Regularity | Regular rhythm |
| Rate (bpm) | > 100 bpm (typically 100–180 bpm) |
| P wave Morphology | Absent, inverted, or after QRS due to retrograde atrial activation |
| Other Notable Features | May result from enhanced AV junction automaticity (e.g., post-surgery, digitalis toxicity, ischemia) |

VENTRICULAR RHYTHMS

Ventricular rhythms are cardiac arrhythmias characterized by the initiation of electrical impulses from the ventricles rather than the sinoatrial (SA) node or atria. In ventricular rhythms, the normal electrical conduction pathway is disrupted, and the ventricular tissue becomes the primary pacemaker.

The atria, AV junction, or both, are unable to initiate an electrical impulse. Purkinje fibers firing rate: 20-40 bpm

Characteristics of ventricular rhythm:

  - No P wave

  - Wide QRS

PREMATURE VENTRICULAR CONTRACTION (PVCs)

ventricular contractions (PVCs) are abnormal heartbeats that originate from the ventricles (the lower chambers of the heart) and occur earlier than expected in the cardiac cycle. They disrupt the heart's regular rhythm, often causing the sensation of a "skipped" or "extra" beat.Ventricular beat occurs before the expected time.There should be a compensatory pause after the premature beats.

- P wave: Absent.

- QRS complex: Wide.

TYPES OF PVCs:

- R on T  phenomenon: T of previous beat & R of the PVC  beat comes close

- Interpolated  PVCs: PVC which is not changing the basic rhythm regularity

| Feature | Interpolated PVCs – Typical Characteristics |
| --- | --- |
| P wave Amplitude (mV / mm) | Present before adjacent normal beats; absent before PVC |
| P wave Width (ms) | Normal in sinus beats; not measurable for PVC |
| PR Interval (ms) | Normal in sinus beats; not measurable for PVC |
| QRS Amplitude (mV / mm) | Normal for sinus beats; PVC usually >1.0 mV / >10 mm |
| QRS Width (ms) | >120 ms for PVC (wide complex) |
| ST Segment Deviation (mV / mm) | Normal in sinus beats; discordant or abnormal in PVC |
| T wave Amplitude (mV / mm) | Normal in sinus beats; discordant/inverted in PVC |
| QT Interval | Appears prolonged due to wide PVC QRS |
| Regularity | Overall rhythm remains regular – PVC is "interpolated" without a compensatory pause |
| Rate (bpm) | Depends on underlying sinus rhythm; PVC does not interrupt the rhythm cycle |
| P wave Morphology | Normal in sinus beats; absent or dissociated in PVC |
| Other Notable Features | PVC occurs between two normal beats without compensatory pause; seen in slower rhythms; may mimic sinus rhythm continuity |

- Unifocal PVCs: Unifocal premature ventricular contractions (PVCs) are abnormal heartbeats originating from a single location within the ventricles, resulting in PVCs with consistent morphology on the ECG

.

| Feature | Characteristics |
| --- | --- |
| P wave Amplitude (mV / mm) | Absent before QRS; may appear after due to AV dissociation |
| P wave Width (ms) | Not applicable (no preceding P wave) |
| PR Interval (ms) | Not measurable (no PR interval before PVC) |
| QRS Amplitude (mV / mm) | >1.0 mV / >10 mm (often high due to ectopic ventricular origin) |
| QRS Width (ms) | >120 ms (wide complex; a key feature of PVCs) |
| ST Segment Deviation (mV / mm) | Discordant with QRS; may show minor secondary repolarization changes |
| T wave Amplitude (mV / mm) | Inverted or discordant in polarity compared to QRS; variable amplitude |
| QT Interval | Appears prolonged due to prolonged QRS, but QTc may be normal |
| Regularity | Irregular due to premature beat; may form regular patterns (e.g., bigeminy, trigeminy) |
| Rate (bpm) | Not a standalone rhythm; PVCs are premature beats within an underlying rhythm |
| P wave Morphology | Absent before PVC; may be retrograde or dissociated |

- Multifocal PVCs: Multifocal premature ventricular contractions (PVCs) are abnormal heartbeats originating from multiple locations within the ventricles, resulting in PVCs with different morphologies on the ECG.

| Feature | Multifocal PVCs – Typical Characteristics |
| --- | --- |
| P wave Amplitude (mV / mm) | Absent before PVCs; may appear after due to AV dissociation |
| P wave Width (ms) | Not applicable for PVCs |
| PR Interval (ms) | Not measurable (no consistent P-QRS relationship for PVCs) |
| QRS Amplitude (mV / mm) | Variable – depends on origin; can range from normal to >1.0 mV / >10 mm |
| QRS Width (ms) | >120 ms (wide complexes); varying morphology indicates multiple foci |
| ST Segment Deviation (mV / mm) | Discordant with each PVC; may differ between morphologies |
| T wave Amplitude (mV / mm) | Discordant to QRS in each PVC; variable polarity |
| QT Interval | Appears prolonged due to wide QRS; interval may vary between different PVCs |
| Regularity | Irregular – PVCs occur unpredictably and from different foci |
| Rate (bpm) | Depends on underlying rhythm; PVCs occur intermittently |
| P wave Morphology | Normal in sinus beats; absent or unrelated in PVCs |
| Other Notable Features | Morphologically different PVCs within the same recording; may suggest increased ventricular irritability or risk for VT |

VENTRICULAR ESCAPE BEATs (VEBs)

  - Occurs after the expected time

  - No compensatory pause after the escape beat

  - P wave: Absent

  - QRS complex: Wide

| Feature | Ventricular Escape Beats – Typical Characteristics |
| --- | --- |
| P wave Amplitude (mV / mm) | Absent before QRS; may appear after due to AV dissociation |
| P wave Width (ms) | Not applicable (no preceding P wave) |
| PR Interval (ms) | Not measurable (no P-QRS relationship) |
| QRS Amplitude (mV / mm) | Often >1.0 mV / >10 mm (but may vary) |
| QRS Width (ms) | >120 ms (wide complex; due to ventricular origin) |
| ST Segment Deviation (mV / mm) | Discordant with QRS; possible minor secondary repolarization changes |
| T wave Amplitude (mV / mm) | Inverted or discordant to QRS polarity |
| QT Interval | May appear prolonged due to wide QRS duration |
| Regularity | Escape beats are typically isolated and occur after a pause (escape interval) |
| Rate (bpm) | Usually 20–40 bpm if escape rhythm is sustained (idioventricular); individual beats are slower |
| P wave Morphology | Absent or dissociated from QRS; no atrial conduction to ventricles |
| Other Notable Features | Occurs late (after a long pause) to rescue ventricular activity; usually due to AV block or sinus arrest; not premature |

TYPES OF PVC TERMINOLOGIES:

- Ventricular Couplets: A ventricular couplet is a cardiac arrhythmia characterized by two consecutive premature ventricular contractions (PVCs) occurring back-to-back. In other words, it involves two successive abnormal heartbeats originating from the ventricles.

| Feature | Typical Characteristics |
| --- | --- |
| P wave Amplitude (mV / mm) | Absent before both QRS complexes; may appear after due to AV dissociation |
| P wave Width (ms) | Not applicable (no preceding P waves) |
| PR Interval (ms) | Not measurable (no PR interval with ectopic ventricular beats) |
| QRS Amplitude (mV / mm) | >1.0 mV / >10 mm (often high; both beats have similar amplitude if unifocal) |
| QRS Width (ms) | >120 ms (both QRS complexes are wide) |
| ST Segment Deviation (mV / mm) | Discordant with QRS; possible secondary repolarization abnormalities |
| T wave Amplitude (mV / mm) | Inverted or discordant in polarity compared to QRS; consistent in unifocal couplets |
| QT Interval | Appears prolonged due to wide QRS complexes |
| Regularity | Irregular overall rhythm due to premature occurrence; couplet itself is regular between two PVCs |
| Rate (bpm) | Not a sustained rhythm; part of underlying rhythm—two premature ventricular beats in sequence |
| P wave Morphology | Absent or dissociated; not associated with QRS of the couplet |
| Other Notable Features | Two consecutive PVCs; identical morphology in unifocal couplets, variable in multifocal |

- Ventricular Triplets/Ventricular Salvo: A ventricular triplet is a cardiac arrhythmia characterized by three consecutive premature ventricular contractions (PVCs) occurring in succession. This means there are three successive abnormal heartbeats originating from the ventricles.

| Feature | Ventricular Triplets – Typical Characteristics |
| --- | --- |
| P wave Amplitude (mV / mm) | Absent before all three QRS complexes; may appear after due to AV dissociation |
| P wave Width (ms) | Not applicable (no preceding P waves) |
| PR Interval (ms) | Not measurable (no PR interval with ventricular beats) |
| QRS Amplitude (mV / mm) | >1.0 mV / >10 mm (typically high; uniform if unifocal) |
| QRS Width (ms) | >120 ms (all three beats are wide complexes) |
| ST Segment Deviation (mV / mm) | Discordant with QRS; possible secondary ST changes |
| T wave Amplitude (mV / mm) | Inverted or discordant with QRS; consistent in unifocal triplets |
| QT Interval | Appears prolonged due to prolonged QRS duration |
| Regularity | Triplet itself is regular; overall rhythm is irregular if isolated triplet |
| Rate (bpm) | Not a sustained rhythm; three consecutive PVCs occurring faster than underlying rhythm |
| P wave Morphology | Absent or dissociated from QRS complexes |
| Other Notable Features | Three consecutive PVCs; considered non-sustained VT if >100 bpm; identical QRS in unifocal triplets |

- Non-Sustained Ventricular Tachycardia (NSVT)

Non-sustained ventricular tachycardia (NSVT) is defined as three or more consecutive ventricular beats occurring at a rate of more than 100 beats per minute, lasting less than 30 seconds, and terminating spontaneously without the need for medical intervention.

### Key Criteria:

- Ventricular origin (wide QRS complexes)

- ≥3 consecutive beats

- Rate: >100 bpm (often >120 bpm)

- Duration: <30 seconds

| Non-Sustained Ventricular Tachycardia (NSVT) |  |
| --- | --- |
| Feature | Typical Values / Characteristics |
| P wave Amplitude (mV / mm) | Absent during NSVT episodes |
| P wave Width (ms) | Not measurable during tachycardia |
| PR Interval (ms) | Not measurable — AV dissociation may be seen |
| QRS Amplitude (mV / mm) | Variable; often large (>2.5 mV) in some leads |
| QRS Width (ms) | Wide: ≥140 ms (usually >160 ms) |
| ST Segment Deviation (mV / mm) | Often distorted by wide QRS; may show secondary ST-T changes |
| T wave Amplitude (mV / mm) | Often opposite in direction to QRS (discordant) |
| QT Interval (ms) | Not reliable during VT; post-conversion QTc should be assessed separately |
| RR Interval (ms) | Regular or slightly irregular; intervals between 200–400 ms (rate 150–300 bpm) |
| Rate (bpm) | 150–250 bpm |
| P wave Morphology | Absent or independent from QRS; AV dissociation may be present |
| Regularity | Usually regular |
| Other Notable Features | ≥3 consecutive wide QRS complexes originating from the ventricles, lasting <30 seconds; terminates spontaneously; can occur in structurally normal or abnormal hearts |

- Ventricular Bigeminy: Every 2nd beat is PVC

| Feature | Ventricular Bigeminy – Typical Characteristics |
| --- | --- |
| P wave Amplitude (mV / mm) | Present before normal beats; absent before PVCs |
| P wave Width (ms) | Normal for sinus beats; not applicable before PVCs |
| PR Interval (ms) | Normal for sinus beats; not measurable before PVCs |
| QRS Amplitude (mV / mm) | Normal for sinus beats; >1.0 mV / >10 mm for PVCs |
| QRS Width (ms) | Normal for sinus beats; >120 ms for PVCs (wide complex) |
| ST Segment Deviation (mV / mm) | May be normal in sinus beats; discordant in PVCs |
| T wave Amplitude (mV / mm) | Normal in sinus beats; discordant in PVCs |
| QT Interval | Normal in sinus beats; appears prolonged in PVCs due to wide QRS |
| Regularity | Regularly irregular – one sinus beat followed by one PVC in a repeating pattern |
| Rate (bpm) | Depends on underlying sinus rhythm; typically regular sinus rate with interspersed PVCs |
| P wave Morphology | Normal in sinus beats; absent or dissociated in PVCs |
| Other Notable Features | Alternating pattern: normal QRS – PVC – normal QRS – PVC; PVCs usually identical (unifocal) |

- Ventricular Trigeminy: Every 3rd beat is PVC

| Feature | Ventricular Trigeminy – Typical Characteristics |
| --- | --- |
| P wave Amplitude (mV / mm) | Present before normal beats; absent before PVCs |
| P wave Width (ms) | Normal for sinus beats; not applicable before PVCs |
| PR Interval (ms) | Normal for sinus beats; not measurable before PVCs |
| QRS Amplitude (mV / mm) | Normal for sinus beats; >1.0 mV / >10 mm for PVCs |
| QRS Width (ms) | Normal for sinus beats; >120 ms for PVCs (wide complex) |
| ST Segment Deviation (mV / mm) | May be normal in sinus beats; discordant in PVCs |
| T wave Amplitude (mV / mm) | Normal in sinus beats; discordant in PVCs |
| QT Interval | Normal in sinus beats; appears prolonged in PVCs due to wide QRS |
| Regularity | Regularly irregular – pattern of two normal beats followed by one PVC |
| Rate (bpm) | Depends on underlying sinus rate; PVCs occur every third beat |
| P wave Morphology | Normal in sinus beats; absent or dissociated in PVCs |
| Other Notable Features | Repeating pattern: normal – normal – PVC – normal – normal – PVC; PVCs typically unifocal |

- Ventricular Quadrigeminy: Every 4th beat is PVC

| Feature | Ventricular Quadrigeminy – Typical Characteristics |
| --- | --- |
| P wave Amplitude (mV / mm) | Present before normal beats; absent before PVCs |
| P wave Width (ms) | Normal for sinus beats; not applicable before PVCs |
| PR Interval (ms) | Normal for sinus beats; not measurable for PVCs |
| QRS Amplitude (mV / mm) | Normal in sinus beats; >1.0 mV / >10 mm for PVCs |
| QRS Width (ms) | Normal in sinus beats; >120 ms for PVCs (wide complex) |
| ST Segment Deviation (mV / mm) | Normal in sinus beats; discordant or altered in PVCs |
| T wave Amplitude (mV / mm) | Normal in sinus beats; inverted or discordant in PVCs |
| QT Interval | Normal in sinus beats; appears prolonged in PVCs due to wide QRS |
| Regularity | Regularly irregular – pattern of three normal beats followed by one PVC |
| Rate (bpm) | Determined by underlying sinus rhythm; PVCs occur every fourth beat |
| P wave Morphology | Normal in sinus beats; absent or dissociated in PVCs |
| Other Notable Features | Repeating pattern: normal – normal – normal – PVC; typically unifocal; full compensatory pause |

IDIOVENTRICULAR RHYTHM (IVR)

- Rate: 20-40bpm

- Rhythm: Regular

- P wave: None

- PR Interval: None

- QRS Complex: Wide(>0.12sec), Bizarre appearance

| Feature | Idioventricular Rhythm – Typical Characteristics |
| --- | --- |
| P wave Amplitude (mV / mm) | Absent or dissociated from QRS |
| P wave Width (ms) | Not applicable (no consistent atrial activity) |
| PR Interval (ms) | Not measurable (no atrial conduction to ventricles) |
| QRS Amplitude (mV / mm) | Often >1.0 mV / >10 mm (variable, depending on origin) |
| QRS Width (ms) | >120 ms (wide QRS due to ventricular origin) |
| ST Segment Deviation (mV / mm) | May be discordant with QRS; secondary ST-T changes possible |
| T wave Amplitude (mV / mm) | Inverted or discordant to QRS polarity |
| QT Interval | Appears prolonged due to wide QRS |
| Regularity | Regular rhythm |
| Rate (bpm) | Typically 20–40 bpm (accelerated idioventricular: 40–100 bpm) |
| P wave Morphology | Absent or retrograde; not related to QRS |
| Other Notable Features | Escape rhythm originating from ventricles; occurs when higher pacemakers fail (e.g., sinus arrest, high-grade AV block); may transition to VT if rate increases |

ACCELERATED IDIOVENTRICULAR RHYTHM (AIVR)

- Rate: 41-100bpm

- Rhythm: Regular

- P wave: None

- PR Interval: None

- QRS Complex: Wide(>0.12sec), Bizarre appearance

| Feature | Accelerated Idioventricular Rhythm (AIVR) – Typical Characteristics |
| --- | --- |
| P wave Amplitude (mV / mm) | Absent or dissociated from QRS |
| P wave Width (ms) | Not applicable |
| PR Interval (ms) | Not measurable (no atrial-to-ventricular conduction) |
| QRS Amplitude (mV / mm) | Often >1.0 mV / >10 mm (may vary with ventricular origin) |
| QRS Width (ms) | >120 ms (wide QRS complex due to ventricular origin) |
| ST Segment Deviation (mV / mm) | May show discordant ST-T changes |
| T wave Amplitude (mV / mm) | Inverted or discordant with QRS |
| QT Interval | May appear prolonged due to widened QRS |
| Regularity | Regular rhythm |
| Rate (bpm) | 40–100 bpm (key distinguishing feature from idioventricular rhythm and VT) |
| P wave Morphology | Absent, dissociated, or retrograde P waves |
| Other Notable Features | Often transient, seen in reperfusion after MI; typically benign and self-limiting; slower than VT, faster than escape rhythm |

AGONAL RHYTHM

  - Ventricular rhythm with HR  less than 20 bpm

  - Also referred as ‘Dying Heart’.

| Feature | Agonal Rhythm – Typical Characteristics |
| --- | --- |
| P wave Amplitude (mV / mm) | Usually absent |
| P wave Width (ms) | Not applicable |
| PR Interval (ms) | Not measurable |
| QRS Amplitude (mV / mm) | Variable, often low amplitude due to poor ventricular activity |
| QRS Width (ms) | Typically wide (>120 ms), but can be variable |
| ST Segment Deviation (mV / mm) | Variable; often indistinct due to low signal quality |
| T wave Amplitude (mV / mm) | Variable; often low or absent |
| QT Interval | Usually difficult to measure |
| Regularity | Highly irregular or almost absent rhythm |
| Rate (bpm) | Extremely slow, typically <20 bpm |
| P wave Morphology | Usually absent or dissociated |
| Other Notable Features | Terminal rhythm seen in cardiac arrest or dying heart; reflects severely depressed ventricular activity; usually precedes asystole |

VENTRICULAR TACHYCARDIA

- Ventricular tachycardia (VT) is a rapid heart rhythm originating from the ventricles (the lower chambers of the heart), characterized by a series of consecutive beats with a rate typically exceeding 100 beats per minute.

- VT can be sustained (lasting longer than 30 seconds) or non-sustained (lasting less than 30 seconds). It often arises due to abnormalities in the heart's electrical conduction system or underlying heart disease.

- Ventricular tachycardia can be life-threatening, especially if it deteriorates into ventricular fibrillation, which can cause sudden cardiac arrest.

TYPES OF VENTRICULAR TACHYCARDIA (VT)

There are 2 types of VT,

  - Monomorphic VT

  - Polymorphic VT

- Monomorphic VT: A ventricular tachycardia with QRS complexes that have a uniform shape and morphology from beat to beat, indicating a single ventricular focus.

| Feature | Monomorphic Ventricular Tachycardia – Typical Characteristics |
| --- | --- |
| P wave Amplitude (mV / mm) | Absent or AV dissociation (P waves may be independent of QRS) |
| P wave Width (ms) | Not applicable |
| PR Interval (ms) | Not measurable (no consistent P-QRS relationship) |
| QRS Amplitude (mV / mm) | Typically >1.0 mV / >10 mm (can be large due to ventricular origin) |
| QRS Width (ms) | >140 ms, often >160 ms (wide and uniform complex) |
| ST Segment Deviation (mV / mm) | Often shows discordant ST-T changes |
| T wave Amplitude (mV / mm) | Opposite polarity to QRS (discordant T wave) |
| QT Interval | Appears prolonged due to wide QRS |
| Regularity | Very regular rhythm (hallmark of monomorphic VT) |
| Rate (bpm) | Typically 100–250 bpm |
| P wave Morphology | P waves may be absent, retrograde, or independent |
| Other Notable Features | All QRS complexes are identical in shape (from a single ectopic focus); may show fusion or capture beats if AV conduction is intact |

- Polymorphic VT: A ventricular tachycardia characterized by QRS complexes that vary in shape, amplitude, and duration, often indicating multiple ventricular foci or changing reentry circuits.

| Feature | Polymorphic Ventricular Tachycardia – Typical Characteristics |
| --- | --- |
| P wave Amplitude (mV / mm) | Absent or not associated with QRS |
| P wave Width (ms) | Not applicable |
| PR Interval (ms) | Not measurable |
| QRS Amplitude (mV / mm) | Variable amplitude (beat-to-beat) |
| QRS Width (ms) | >120 ms (all wide, but shape and axis vary) |
| ST Segment Deviation (mV / mm) | Often obscured; may show discordant ST-T changes |
| T wave Amplitude (mV / mm) | Often not clearly distinguishable |
| QT Interval | Often prolonged prior to onset if torsades de pointes (a subtype) |
| Regularity | Irregular – varying QRS morphology and axis |
| Rate (bpm) | Typically 150–250 bpm |
| P wave Morphology | Absent or dissociated from QRS |
| Other Notable Features | QRS complexes change in shape, size, and axis; seen in acute ischemia, electrolyte imbalance, prolonged QT (Torsades), or genetic syndromes like Brugada/LQTS |

TORSADES DE POINTES (TDP)

- Torsades de pointes (TdP) is a specific form of polymorphic ventricular tachycardia characterized by a unique twisting or "twisting of the points" appearance on the electrocardiogram (ECG).

- It is typically associated with a prolonged QT interval, which predisposes individuals to the development of TdP.

- TDP can be triggered by various factors, including medications that prolong the QT interval, electrolyte abnormalities, and certain cardiac conditions.

- It has the potential to degenerate into ventricular fibrillation and cause sudden cardiac arrest if not promptly treated.

| Feature | Torsades de Pointes (TdP) – Typical Characteristics |
| --- | --- |
| P wave Amplitude (mV / mm) | Absent during episodes |
| P wave Width (ms) | Not applicable |
| PR Interval (ms) | Not measurable |
| QRS Amplitude (mV / mm) | Variable amplitude, often waxing and waning |
| QRS Width (ms) | >120 ms, with changing morphology and axis |
| ST Segment Deviation (mV / mm) | Often obscured by rapid ventricular activity |
| T wave Amplitude (mV / mm) | Difficult to evaluate during episode; T wave abnormalities (e.g., notched, prolonged QT) often precede onset |
| QT Interval | Prolonged QT (>450 ms in men, >470 ms in women) before TdP onset |
| Regularity | Irregular; rhythm varies with the changing QRS morphology |
| Rate (bpm) | Typically 200–250 bpm |
| P wave Morphology | Absent or dissociated |
| Other Notable Features | Twisting of QRS complexes around the isoelectric line (classic TdP appearance); occurs in the setting of prolonged QT, hypokalemia, hypomagnesemia, medications, or genetic syndromes (e.g., LQTS) |

VENTRICULAR FIBRILLATION (VF)

- Ventricular fibrillation (VF) is a life-threatening cardiac arrhythmia characterized by rapid and chaotic electrical activity in the ventricles (the lower chambers of the heart).

- In VF, the heart's normal electrical impulses become disorganized, leading to ineffective and irregular ventricular contractions. This results in the loss of coordinated pumping function, causing the heart to quiver instead of contracting effectively.

- VF is a medical emergency that can lead to sudden cardiac arrest and death within minutes if not treated promptly with cardiopulmonary resuscitation (CPR) and defibrillation

| Feature | Ventricular Fibrillation (VF) – Typical Characteristics |
| --- | --- |
| P wave Amplitude (mV / mm) | Absent |
| P wave Width (ms) | Not applicable |
| PR Interval (ms) | Not measurable |
| QRS Amplitude (mV / mm) | No organized QRS complexes; chaotic waveform with variable amplitude |
| QRS Width (ms) | Not measurable |
| ST Segment Deviation (mV / mm) | Not discernible |
| T wave Amplitude (mV / mm) | Not discernible |
| QT Interval | Not measurable |
| Regularity | Totally irregular – no repeating pattern |
| Rate (bpm) | Not measurable; ventricular rate is disorganized and rapid (>300 bpm in fine VF) |
| P wave Morphology | Absent |
| Other Notable Features | Chaotic, disorganized electrical activity; no effective cardiac output; can appear as coarse (high amplitude) or fine (low amplitude) VF |

VENTRICULAR STAND STILL (VSS)

- "Ventricular standstill," also known as "ventricular asystole," refers to a condition in which the ventricles (the lower chambers of the heart) cease to contract and exhibit no electrical activity.

- This results in the absence of a detectable heartbeat on electrocardiogram (ECG) monitoring. Ventricular standstill is a critical medical emergency and is often associated with a complete cessation of cardiac output.

- Immediate intervention, such as cardiopulmonary resuscitation (CPR) and advanced cardiac life support (ACLS) measures, is required to restore cardiac function and circulation.

| Feature | Ventricular Standstill – Typical Characteristics |
| --- | --- |
| P wave Amplitude (mV / mm) | Present, often normal amplitude |
| P wave Width (ms) | Normal (typically 80–110 ms) |
| PR Interval (ms) | Not measurable – no ventricular response |
| QRS Amplitude (mV / mm) | Absent – no QRS complexes |
| QRS Width (ms) | Not measurable |
| ST Segment Deviation (mV / mm) | Not discernible |
| T wave Amplitude (mV / mm) | Not discernible |
| QT Interval | Not measurable |
| Regularity | Atrial rhythm may be regular, but ventricular activity is absent |
| Rate (bpm) | Atrial rate may be normal or slow (e.g., 40–60 bpm); ventricular rate is 0 bpm |
| P wave Morphology | Normal sinus P waves |
| Other Notable Features | Indicates complete failure of ventricular activity; SA node may be functional, but no AV conduction or ventricular escape; considered a form of asystole |

PULSELESS ELECTRICAL ACTIVITY (PEA):

  - Only electrical/ chemical  activities are taking  place, so ECG tracings  seen

  - No mechanical activity,  No contraction, No  pulse, No cardiac output

| Feature | Pulseless Electrical Activity (PEA) – Typical Characteristics |
| --- | --- |
| P wave Amplitude (mV / mm) | Variable – may be present or absent depending on the underlying rhythm |
| P wave Width (ms) | Normal if present |
| PR Interval (ms) | May be measurable if P waves are present; often normal |
| QRS Amplitude (mV / mm) | Normal or low depending on underlying rhythm |
| QRS Width (ms) | Normal or wide – can be narrow (e.g., sinus rhythm) or wide (e.g., idioventricular) |
| ST Segment Deviation (mV / mm) | Depends on underlying cause; may show signs of ischemia or be normal |
| T wave Amplitude (mV / mm) | Variable |
| QT Interval | Not typically significant in diagnosis |
| Regularity | May be regular or irregular, depending on rhythm |
| Rate (bpm) | Any rate – there is electrical activity, but no palpable pulse |
| P wave Morphology | Normal or abnormal depending on underlying rhythm |
| Other Notable Features | Electrical activity present on ECG without effective mechanical cardiac output (no pulse); associated with reversible causes – H’s and T’s (e.g., hypoxia, hypovolemia, tamponade, tension pneumothorax, thrombosis, toxins) |

ASYSTOLE:

  - No electrical/ chemical  activity

  - No mechanical activity,  No contraction, No  pulse, No cardiac output

| Feature | Asystole – Typical Characteristics |
| --- | --- |
| P wave Amplitude (mV / mm) | Absent (occasionally isolated, non-conducted P waves may appear briefly) |
| P wave Width (ms) | Not measurable |
| PR Interval (ms) | Not measurable |
| QRS Amplitude (mV / mm) | Absent – no QRS complexes |
| QRS Width (ms) | Not measurable |
| ST Segment Deviation (mV / mm) | Not discernible |
| T wave Amplitude (mV / mm) | Not discernible |
| QT Interval | Not measurable |
| Regularity | None – flat or nearly flat isoelectric line |
| Rate (bpm) | 0 bpm (no electrical or mechanical activity) |
| P wave Morphology | Absent, or if present, no associated QRS complexes |
| Other Notable Features | Represents complete cessation of electrical activity in the heart; no pulse, no cardiac output; most extreme form of cardiac arrest |

HEART BLOCKS

Heart block refers to partial delays or complete interruptions in the cardiac conduction pathway between the atria (the upper chambers of the heart) and ventricles (the lower chambers).

The degree of blockage determines the type and classification of heart block. This condition disrupts the normal electrical signals that coordinate heartbeats, leading to abnormalities in heart rhythm and potential symptoms such as dizziness, fainting, or chest pain.

Heart block can vary in severity, ranging from mild to life-threatening, and may require medical intervention depending on its impact on cardiac function and symptoms experienced by the individual.

Common causes:

  - Ischemia

  - Myocardial necrosis

  - Degenerative disease of the conduction system

  - Congenital anomalies

  - Drugs (especially digitalis preparations)

Types of AV Blocks:

- First-Degree AV Block

- Second-Degree AV Block:

  - Type I (Wenckebach or Mobitz I)

  - Type II (Mobitz II)

- Third-Degree AV Block (Complete Heart Block)

1st Degree AV Block:

First-degree atrioventricular block (AVB) is a type of conduction abnormality in the heart characterized by a delay in the conduction of electrical impulses from the atria to the ventricles. In this condition, every atrial impulse is conducted to the ventricles, but the conduction time is prolonged, leading to a lengthened PR interval on the electrocardiogram (ECG).

Characteristics:

- Abnormally prolonged PR interval

- PR interval > 0.20 seconds and constant

Significance:

- Usually not significant

- Underlying cause should be corrected

- Usually does not produce symptoms

| Feature | First-Degree AV Block – Typical Characteristics |
| --- | --- |
| P wave Amplitude (mV / mm) | Normal |
| P wave Width (ms) | Normal (80–110 ms) |
| PR Interval (ms) | Prolonged – >200 ms (≥5 small boxes on ECG) |
| QRS Amplitude (mV / mm) | Normal |
| QRS Width (ms) | Normal (<120 ms) unless coexisting bundle branch block |
| ST Segment Deviation (mV / mm) | Normal |
| T wave Amplitude (mV / mm) | Normal |
| QT Interval | Normal |
| Regularity | Regular atrial and ventricular rhythm |
| Rate (bpm) | Depends on underlying sinus rhythm |
| P wave Morphology | Normal |
| Other Notable Features | Each P wave is followed by a QRS, but with a consistently prolonged PR interval; typically benign, but may indicate early conduction system disease or be medication-related |

2nd Degree AV Block Type 1:

- Arrhythmia in which there is a progressive delay following each P wave in the conduction of electrical impulses through the AV node until conduction is completely blocked

- Most commonly represents defective conduction through the AV node or occasionally in the Bundle of HIS or bundle branches

- Second-degree type 1 atrioventricular (AV) block, also known as Wenckebach or Mobitz type 1 AV block, is a conduction abnormality in the heart characterized by progressive prolongation of the PR interval followed by a dropped beat (missing QRS complex) on the electrocardiogram (ECG).

- In this type of AV block, the delay in conduction through the AV node gradually increases with each successive beat until an impulse fails to conduct to the ventricles, resulting in a dropped beat. After the dropped beat, the cycle repeats, with subsequent PR intervals shortening and then lengthening again.

- Second-degree type 1 AV block is often benign and may be asymptomatic, but it can also occur transiently in healthy individuals during sleep or with certain medications. However, it can also be associated with underlying heart disease or conditions affecting the conduction system of the heart, warranting further evaluation and monitoring.

Characteristics

- Progressive lengthening of the PR interval until a QRS complex fails to appear after a P wave

- Repetitive

- Also called as Mobitz type I or Wenckebach

Significance

- Usually transient and reversible

- Usually produces few symptoms

- May progress to a higher degree block

- Will respond to atropine if rate requires treatment

| Feature | Second-Degree AV Block – Type I (Mobitz I / Wenckebach) |
| --- | --- |
| P wave Amplitude (mV / mm) | Normal |
| P wave Width (ms) | Normal (80–110 ms) |
| PR Interval (ms) | Progressively lengthens with each beat until a QRS is dropped |
| QRS Amplitude (mV / mm) | Normal |
| QRS Width (ms) | Normal (<120 ms) unless coexisting bundle branch block |
| ST Segment Deviation (mV / mm) | Normal |
| T wave Amplitude (mV / mm) | Normal |
| QT Interval | Normal |
| Regularity | Regularly irregular – grouped beating pattern |
| Rate (bpm) | Atrial rate is regular; ventricular rate is slower due to dropped beats |
| P wave Morphology | Normal |
| Other Notable Features | Characterized by progressive PR prolongation followed by a non-conducted P wave (dropped QRS); typically benign, often seen in athletes, during sleep, or with increased vagal tone |

2nd Degree AV Block Type 2:

- An arrhythmia in which a complete block of conduction of the electrical impulses occurs in one bundle branch and an intermittent block in the other.

- Second-degree type 2 atrioventricular (AV) block, also known as Mobitz type 2 AV block, is a conduction abnormality in the heart characterized by intermittent failure of atrial impulses to conduct to the ventricles without progressive prolongation of the PR interval on the electrocardiogram (ECG).

- In this type of AV block, some atrial impulses fail to conduct to the ventricles, resulting in dropped beats (missing QRS complexes) on the ECG.

- Unlike type 1 AV block, the PR interval remains constant before the dropped beats.

- Second-degree type 2 AV block is considered more serious than type 1 AV block and may indicate a higher risk of progression to complete heart block.

- It is often associated with underlying heart disease or conditions affecting the conduction system of the heart and may require close monitoring and management

- Characteristics

An AV block with a regularly or irregularly absent QRS complex, commonly producing an AV conduction ratio of 4:3 or 3:2

Conduction ratio:

- Every other P wave conducts QRS - 2:1 conduction

- Every four P waves conducted there are 3 QRS conducted - 4:3 conduction

- Variable conduction will result in irregular ventricular (R-R) rhythm

- Significance

- More serious

- Often progresses to 3rd degree

- Sx consistent with symptomatic bradycardia

- Atropine is usually not effective

- Be prepared for TCP

| Feature | Second-Degree AV Block – Type II (Mobitz II) |
| --- | --- |
| P wave Amplitude (mV / mm) | Normal |
| P wave Width (ms) | Normal (80–110 ms) |
| PR Interval (ms) | Constant (normal or slightly prolonged) before conducted beats |
| QRS Amplitude (mV / mm) | Normal or low amplitude, especially if associated with bundle branch block |
| QRS Width (ms) | Often wide (>120 ms), especially with coexisting bundle branch block |
| ST Segment Deviation (mV / mm) | Usually normal |
| T wave Amplitude (mV / mm) | Normal |
| QT Interval | May be prolonged if QRS is wide |
| Regularity | Irregular – due to intermittent dropped QRS complexes |
| Rate (bpm) | Atrial rate is regular; ventricular rate is slower depending on the number of dropped beats |
| P wave Morphology | Normal |
| Other Notable Features | Sudden, unpredictable dropped QRS with no PR interval prolongation; can progress to complete heart block; usually indicates infranodal disease (His–Purkinje); more dangerous than Mobitz I |

3rd Degree AV Block:

- Complete absence of conduction of the electrical impulses through the AV node, bundle of HIS, or bundle branches, Characterized by independent beating of the atria and ventricles.

- Third-degree atrioventricular (AV) block, also known as complete heart block, is a severe conduction abnormality in the heart characterized by complete dissociation between the atria and ventricles.

- In this type of AV block, no atrial impulses are conducted to the ventricles, resulting in independent and asynchronous beating of the atria and ventricles. As a result, the atria and ventricles beat at their intrinsic rates, leading to a slow ventricular escape rhythm. On the electrocardiogram (ECG), third-degree AV block is characterized by dissociated P waves (atrial activity) and QRS complexes (ventricular activity), with no consistent relationship between them.

- Third-degree AV block is a medical emergency that can result in severe symptoms, such as syncope (fainting), dizziness, shortness of breath, and even sudden cardiac arrest. Prompt intervention, such as implantation of a pacemaker, is usually necessary to restore normal heart rhythm and prevent complications.

Significance

- Sx consistent with symptomatic bradycardia, more ominous

- If Av junction or ventricular escape does not take over following a sudden onset of #rd degree ventricular asystole will occur

- TCP3 rd degree with narrow QRS may respond to atropine

| Feature | Third-Degree AV Block – Typical Characteristics |
| --- | --- |
| P wave Amplitude (mV / mm) | Normal |
| P wave Width (ms) | Normal (80–110 ms) |
| PR Interval (ms) | Variable and unrelated to QRS (no consistent PR intervals) – AV dissociation |
| QRS Amplitude (mV / mm) | May be normal or low, depending on the escape focus |
| QRS Width (ms) | Wide if escape rhythm is ventricular; narrow if escape is junctional |
| ST Segment Deviation (mV / mm) | Usually normal, unless ischemia or other condition present |
| T wave Amplitude (mV / mm) | Normal or inverted |
| QT Interval | May appear prolonged if QRS is wide |
| Regularity | Atrial and ventricular rhythms are regular, but independent of each other |
| Rate (bpm) | Atrial rate: normal; Ventricular rate: 30–60 bpm (junctional escape) or <40 bpm (ventricular) |
| P wave Morphology | Normal, but not related to QRS complexes |
| Other Notable Features | Complete AV dissociation: no atrial impulses are conducted to the ventricles; QRS is from escape focus; may present with syncope (Stokes-Adams attacks), dizziness, hypotension |

Atrial Tachycardia with Variable block (ATV):

Atrial tachycardia is a type of arrhythmia — an irregular heart rhythm — that causes the upper chambers (atria) of your heart to beat faster than normal. They can beat 100 to 250 beats per minute. Atrial tachycardia is a type of SVT or supraventricular tachycardia. Multiples P waves (all the same) not followed by QRS complexes.

PACEMAKER RHYTHMS:

Pacemaker rhythm refers to the cardiac rhythm generated and maintained by an artificial pacemaker device implanted in the heart. These devices are typically implanted in individuals with certain types of heart rhythm disorders or conduction abnormalities that compromise the heart's ability to maintain a regular heartbeat.

Pacemaker rhythms can be identified on an electrocardiogram (ECG) by characteristic patterns, including pacing spikes (indicating when the pacemaker delivers an electrical impulse) and paced QRS complexes (representing ventricular depolarization initiated by the pacemaker).

Indications for Pacemaker

- Brady arrhythmias

- Congenital complete heart block (CHB)- normal or abnormal heart structure

- Trauma- surgical or other

- Slow sinus or junctional rhythm

- Suppression of ectopy

- Permanent pacer malfunction

- Drugs, electrolyte imbalances

- Sick Sinus Syndrome

Modes of Pacing

  - Atrial pacing

- Intact AV conduction system required

  - Ventricular pacing

- Loss of atrial kick

- Discordant ventricular contractions

- Sustains cardiac output

  - Atrial/Ventricular pacing

- Natural pacing

- Atrial-ventricular synchrony

PACEMAKER MALFUNCTIONS:

Failure to sense and failure to capture are two distinct issues related to pacemaker function:

- Failure to capture: This occurs when the pacemaker delivers electrical impulses to the heart, but the heart does not respond by contracting as expected. The pacing stimuli may be delivered at an inadequate energy level, or there may be issues with the pacing lead or myocardial refractoriness, preventing the heart from depolarizing in response to the pacing stimuli.

Examples of FTC:

- Failure to sense: This occurs when the pacemaker fails to detect intrinsic electrical activity in the heart. In other words, the pacemaker does not recognize the heart's natural electrical signals, such as atrial or ventricular depolarizations. As a result, the pacemaker may deliver unnecessary pacing stimuli or fail to adjust pacing rates in response to changes in the heart's intrinsic rhythm.

Examples of FTS:

| Feature | Failure to Capture (FTC) | Failure to Sense (FTS) |
| --- | --- | --- |
| Pacing spike present | ✅ Yes | ✅ Yes |
| QRS complex after spike | ❌ Absent (no ventricular capture) | ✅ Present (but spike may fall in inappropriate timing) |
| P wave after atrial spike | ❌ Absent (in atrial FTC) | ✅ Present (if intrinsic P not sensed) |
| Spike-to-QRS interval | Present but no following QRS | Present but misaligned (e.g., during T wave or refractory period) |
| QRS morphology | Normal intrinsic OR bradycardic rhythm | Varies – may be paced or intrinsic depending on device behavior |
| Atrial or ventricular rate | Bradycardia may result | Irregular pacing or inappropriate pacing during intrinsic rhythm |
| Pacing rate | Normal (as set by pacemaker) | Uncoordinated or too frequent (pacing despite intrinsic activity) |
| Underlying rhythm | Present if capture fails and intrinsic rhythm is preserved | Present but not sensed, so pacing occurs inappropriately |
| Pacing spike timing | On time, but ineffective | Not inhibited by intrinsic activity (misaligned) |
| Key Indicator | Spike not followed by expected depolarization (no QRS/P) | Spike occurs despite intrinsic activity (e.g., on T wave) |
| Risk | Severe bradycardia, syncope | R-on-T phenomenon → ventricular arrhythmia risk |

Sinus Pause

A sinus pause is a transient failure of the sinus node to generate an impulse, resulting in a temporary absence of P waves and their associated QRS complexes.

#### Key Characteristics:

- Duration of the pause is less than 3 seconds.

- No atrial (P wave) activity during the pause.

Sinus Arrest

Sinus arrest refers to a prolonged failure of the sinus node to initiate an impulse, causing a pause ≥3 seconds in the ECG without P waves or QRS complexes.

#### Key Characteristics:

- Longer than 3 seconds

- Not followed by a compensatory mechanism unless an escape rhythm (junctional or ventricular) takes over.

| Feature | Sinus Pause | Sinus Arrest |
| --- | --- | --- |
| P wave Amplitude (mV / mm) | Normal when present: 0.1–0.3 mV (1–3 mm) | Normal when present: 0.1–0.3 mV (1–3 mm) |
| P wave Width (ms) | Normal when present: ≤120 ms (usually 80–110 ms) | Normal when present: ≤120 ms (usually 80–110 ms) |
| PR Interval (ms) | Normal: 120–200 ms when measurable | Normal: 120–200 ms when measurable |
| QRS Amplitude (mV / mm) | Normal: 0.5–2.5 mV (5–25 mm) | Normal: 0.5–2.5 mV (5–25 mm) |
| QRS Width (ms) | Normal: ≤100 ms (usually 80–100 ms) | Normal: ≤100 ms (usually 80–100 ms) |
| ST Segment Deviation (mV / mm) | Isoelectric when present (<1 mm or <0.1 mV) | Isoelectric when present (<1 mm or <0.1 mV) |
| T wave Amplitude (mV / mm) | Normal when present: 0.1–0.5 mV (1–5 mm), upright in most leads | Normal when present: 0.1–0.5 mV (1–5 mm), upright in most leads |
| QT Interval | Normal: 350–440 ms (QTc <440 ms men, <460 ms women) when measurable | Normal: 350–440 ms (QTc <440 ms men, <460 ms women) when measurable |
| Regularity | Interrupted (brief pause), then returns to baseline sinus rhythm | Markedly irregular (pause ≥3 sec), rhythm may not resume without escape beat |
| Rate (bpm) | Variable; slows during pause, otherwise 60–100 bpm | Variable; absent during arrest, may resume with escape beat or remain absent |
| P wave Morphology | Normal before and after pause; upright in I, II, aVF; inverted in aVR | Normal before arrest; absent during arrest |
| Other Notable Features | Pause <3 seconds; no P or QRS during pause; no escape beat unless prolonged | Pause ≥3 seconds; no P or QRS; escape rhythm may occur if arrest is prolonged |

WPW Syndrome

Wolff-Parkinson-White (WPW) syndrome is a pre-excitation syndrome characterized by the presence of an accessory conduction pathway (usually the Bundle of Kent) that connects the atria and ventricles, bypassing the AV node, resulting in early ventricular depolarization and the potential for reentrant tachyarrhythmias.

To diagnose WPW syndrome on an ECG, the following three classic features must be present:

- Short PR interval (< 120 ms)

- Delta wave (slurred upstroke of the QRS complex)

- Prolonged QRS duration (> 120 ms)

| Feature | Characteristics |
| --- | --- |
| P wave Amplitude (mV / mm) | Normal: 0.1–0.3 mV (1–3 mm) |
| P wave Width (ms) | Normal: ≤120 ms (usually 80–110 ms) |
| PR Interval (ms) | Shortened: <120 ms (due to early ventricular activation via accessory pathway) |
| QRS Amplitude (mV / mm) | Variable: typically, normal to high (≥0.5 mV; 5 mm or more) |
| QRS Width (ms) | Prolonged: >120 ms (due to fusion of normal and pre-excited ventricular conduction) |
| ST Segment Deviation (mV / mm) | May be discordant to QRS; nonspecific ST-T changes are common |
| T wave Amplitude (mV / mm) | May be inverted in leads opposite the delta wave (secondary repolarization changes) |
| QT Interval | Normal to mildly shortened; interpret with caution due to distorted QRS |
| Regularity | Regular when in sinus rhythm; may be irregular during tachyarrhythmias |
| Rate (bpm) | Normal (60–100 bpm) in resting sinus rhythm; may be 150–250 bpm in tachyarrhythmias |
| P wave Morphology | Normal (upright in I, II, aVF; inverted in aVR); may be hidden during tachycardia |
| Other Notable Features | Delta wave (slurred upstroke of QRS) - Short PR interval - Wide QRS - May present with AVRT or atrial fibrillation with rapid ventricular response |

Wandering Atrial Pacemaker (WAP)

Wandering atrial pacemaker is a benign atrial rhythm characterized by at least three different P wave morphologies on ECG, with heart rate <100 bpm, due to shifting of the atrial pacemaker site between the SA node, atrial tissue, and/or AV junction.

- P wave morphology: Varies (≥3 distinct forms)

- Heart rate: <100 bpm

- Rhythm: Slightly irregular

| Feature | Typical Values / Characteristics |
| --- | --- |
| P wave Amplitude (mV / mm) | Varies between 0.05–0.25 mV |
| P wave Width (ms) | Varies between 80–110 ms, but can be slightly wider or narrower |
| PR Interval (ms) | Varies with each beat; typically within 120–200 ms, but may differ depending on site of origin |
| QRS Amplitude (mV / mm) | Normal: 0.5–2.5 mV |
| QRS Width (ms) | Normal: <120 ms |
| ST Segment Deviation (mV / mm) | Usually normal |
| T wave Amplitude (mV / mm) | Normal: typically 0.1–0.3 mV |
| QT Interval (ms) | Normal: QTc <450 ms |
| RR Interval (ms) | Mild variability due to changing pacemaker sites |
| Rate (bpm) | <100 bpm (distinguishes it from Multifocal Atrial Tachycardia, which is >100 bpm) |
| P wave Morphology | At least 3 distinct P wave morphologies within a single ECG strip |
| Regularity | Irregular — due to variable pacemaker origin |
| Other Notable Features | Beat-to-beat variability in P wave shape, PR interval, and PP interval; normal QRS; benign rhythm |

Multifocal Atrial Tachycardia (MAT)

Multifocal atrial tachycardia is an irregular supraventricular tachycardia with ≥3 distinct P wave morphologies, and a heart rate >100 bpm, typically associated with underlying pulmonary disease (e.g., COPD).

- P wave morphology: Varies (≥3 distinct forms)

- Heart rate: >100 bpm

- Rhythm: Irregularly irregular

| Feature | Typical Values / Characteristics |
| --- | --- |
| P wave Amplitude (mV / mm) | Varies beat to beat: 0.05–0.25 mV |
| P wave Width (ms) | Varies: typically, 80–120 ms |
| PR Interval (ms) | Varies: 100–200 ms, depending on atrial focus |
| QRS Amplitude (mV / mm) | Normal: 0.5–2.5 mV |
| QRS Width (ms) | Normal: <120 ms |
| ST Segment Deviation (mV / mm) | Usually normal or minor deviations; <1 mm possible |
| T wave Amplitude (mV / mm) | Normal or slightly altered: typically, 0.1–0.3 mV |
| QT Interval (ms) | Normal; QTc typically <450 ms |
| RR Interval (ms) | Irregular — varies due to changing atrial pacemaker activity |
| Rate (bpm) | >100 bpm (typically 100–150 bpm) |
| P wave Morphology | ≥3 different P wave morphologies within one ECG strip |
| Regularity | Irregularly irregular |
| Other Notable Features | Multiple ectopic atrial foci; irregular rhythm with varying P waves, PR intervals, and PP intervals; common in elderly or patients with COPD or hypoxia |

Artifacts

Artifacts are non-cardiac signals that interfere with ECG interpretation. They can distort morphology, mimic arrhythmias, or hide key features.

Types of ECG Artifacts

1. Motion Artifact

Electrical interference in the ECG signal caused by physical movement of the patient or displacement of electrodes relative to the skin.

- Sudden, large baseline shifts.

- Spike-like deflections not aligned with QRS.

- Lead I may show interference while Lead II stays normal (or vice versa).

- Example: Walking, shifting position while recording.

### Muscle Artifact (EMG Noise)

High-frequency electrical activity from skeletal muscles (electromyogram) that contaminates the ECG signal.

- High-frequency, low-amplitude noise overlaid on QRS and T waves.

- Appears as "fuzz" or random jagged lines.

- Common during shivering, talking, or coughing.

- More prominent in limb leads.

### Electrode Contact Noise

Signal corruption resulting from poor or intermittent contact between the ECG electrodes and the skin.

- Flatlines or abrupt dropout of signal.

- Sudden large transient noise when connection re-establishes.

- One channel may go flat while the other records normally.

### Power Line Interference (50/60 Hz)

Electrical noise introduced from surrounding devices or infrastructure operating at mains frequency (typically 50 or 60 Hz).

- Sinusoidal wave superimposed over the ECG baseline.

- Constant amplitude and frequency, often obscuring small deflections.

- Identical pattern on both leads.

### Baseline Wander

Low-frequency drift in the ECG baseline, often due to respiration, electrode impedance changes, or body movement.

- Smooth, wave-like slow shifts in baseline over many beats.

- Caused by respiration or slow movements.

- Preserved QRS morphology but distorted amplitude measurements.

| Artifact Type | Cause | ECG Effect | Frequency Range | Key Characteristics |
| --- | --- | --- | --- | --- |
| Motion Artifacts | Movement of electrodes or patient motion | Sudden baseline shifts, transient noise spikes | 0.5–10 Hz | Baseline wander, inconsistent across leads |
| Muscle (EMG) Artifacts | Muscle contraction (esp. chest wall) | Rapid, low-amplitude noise over QRS/T | 20–200 Hz | Random, spiky, high-frequency, non-rhythmic |
| Electrode Contact Noise | Loose or dry electrodes | Loss of signal, 50/60 Hz noise, flatline sections | Broad spectrum or 50/60 Hz | Flat segments, large baseline shifts, dropout |
| Power Line Interference | Nearby electrical devices | Continuous 50/60 Hz sinusoidal signal | Exactly 50 or 60 Hz | Constant amplitude sine wave |
| Baseline Wander | Respiration or patient movement | Slow oscillation of baseline | <0.5 Hz | Smooth wave drift, low frequency |
| Cable or Lead Noise | Movement or damage to leads | Sharp deflections or periodic noise bursts | Variable | Abrupt changes, repeating interference pattern |

| Artifact Type | Typical Amplitude | Lead Dependency | Comments |
| --- | --- | --- | --- |
| Motion artifact | 0.1 – 1.0 mV | Usually asymmetrical | Mimics arrhythmia, often irregular |
| EMG artifact | <0.2 mV | All leads | Spiky, hard to filter without distortion |
| Electrode noise | >1.0 mV | Often in one lead | Sudden onset and termination |
| Powerline interference | ~0.05 – 0.2 mV | All leads | Can be reduced by notch filters |
| Baseline wander | 0.1 – 0.5 mV | All leads | Harmful to heart rate estimation |